Determining the HUMAC NORM Servo and Motor Type

# Procedure

## If you have an original CYBEX NORM which was not upgraded with a HUMAC, you have a DM servo and Rockwell Motor. Proceed to Step 2 “Your System”.

## If you have a NORM with a HUMAC System, continue with the following steps.

## If you are running the current HUMAC2015 program, from the HELP menu, select ABOUT. The HUMAC will display your interface type (32 or 32/64-bit), and the System, Motor, and Servo.

Figure 1 32-bit Interface

Figure 2 32/64-bit Interface

## If you do not see the Interface and Motor, Servo information, or are running a previous version of the HUMAC, continue with the following steps.

## Exit the HUMAC Program.

## If you have a 32-bit Interface.

Figure 3 32-bit Interface

### From the WINDOWS TASK BAR, select START, ALL PROGRAMS, HUMAC20??, UTILITIES, DIAGNOSTICS where HUMAC20?? Is your HUMAC version, e.g. HUMAC2004 or HUMAC20015.

### From the STARTUP SCREEN, enter CSMIDIAG as the password and click the PROCEED WITH CAUTION button. The HUMAC will display the USB-DSP-PLD MONITOR form.

### The USB-DSP-PLD MONITOR form will display the current Machine, Servo, Motor and Rev settings on the right-hand side of the screen.

## If you have the HUMAC 32/64-bit Interface

Figure 4 32/64-bit Interface

### From the WINDOWS TASK BAR, select START, ALL PROGRAMS, HUMAC2015, UTILITIES, DIAGNOSTICS64.

### From the STARTUP SCREEN, enter CSMIDIAG as the password and click the PROCEED WITH CAUTION button. The HUMAC will display the HUMAC64 DIAGNOSTICS form.

### The DIAGNOSTICS form will display the current MACHINE, SERVO, MOTOR and REV settings on the right-hand side of the screen.

## You can also determine the Motor type from the HUMAC Interface:

| Rockwell Motor |  | Getty Motor |  |
| --- | --- | --- | --- |

## Click the Exit button to exit the diagnostics.

# Your System

## Check the setting for your system.

| Item | Type |
| --- | --- |
| HUMAC Interface | 32-bit, 64-bit |
| Servo | DM-30, DM-75, Ultra3000 |
| Motor | Rockwell, Getty |

## Complete the contact information below and FAX (781-297-2039) or e-mail (orders@csmisolutions.com) this page to CSMi.

| Facility Name: |  |
| --- | --- |
| Street: |  |
| City, State: |  |
| Country: |  |
| Your Name: |  |
| Phone: |  |
| E-Mail: |  |

Update the document with this information.

# HUMAC Interface

## Windows Compatibility

| 32-bit Windows Compatible Interface | 32 and 64-bit Windows Compatible Interface |
| --- | --- |

## Servo Enclosure Compatibility

| Compatible with DM Servo Enclosure Cabling. Note: Enclosure May Contain DM or Ultra3000 Servo with Upgrade Kit | Compatible with Ultra3000 Enclosure Cabling and Ultra3000 Servo |
| --- | --- |

# Servo Amplifiers

The following servo amplifiers have been used with NORM systems.

| Rockwell DM-30 | Rockwell DM-75 | Rockwell Ultra3000 | AMC C060A400 | AMC C060L400 |
| --- | --- | --- | --- | --- |
| Approx 3” wide | Approx 6” wide |  | Three large green connectors on the side | Two large green connectors on the side |

# Motors

The following motors have been used with NORM systems.

| Rockwell motor. - Green or Blue paint. |  |
| --- | --- |
| Getty motor. - Black paint. - Cooling fins are not visible at the back of the motor. - Manufacturer’s label is not visible. |  |
| 507202 motor. - Black paint. - Cooling fins are visible at the back of the motor. - Manufacturer’s label is visible. |  |